## Slide 1
UNANOFABTOOLS · TOOL
DAT Tools
DATfixer & DATgrapher — reading the Denton 635's logs.
A plain-English walkthrough.
Faith
Server Developer & Maintainer
Utah Nanofab · University of Utah
May 2026
Notes: This is a short, self-contained session on two small command-line utilities: DATfixer and DATgrapher. They exist to make the Denton 635 sputter system's log files readable. No coding background needed. By the end you'll know what problem they solve, how to run them, and which one to reach for. It's a good example of a 'small but real' tool in the repository.

## Slide 2
The problem: unreadable machine logs
UNanofabTools — DAT Tools
2
The Denton 635 writes its logs in its own binary format.
Every run, the Denton 635 saves a log file ending in .DAT.
That file is 'binary' — raw bytes, not letters; open it and you see gibberish.
The useful data (pressures, timestamps, events) is in there, just encoded.
We need a way to translate it into something a person can read.
Notes: Set up the why. The Denton 635 is an automated sputter-deposition system. Each run it writes a .DAT log in its own binary format — meaning it stores numbers and control codes as raw bytes, not readable text. Open it in Notepad and it's a jumble. But the real data is in there: chamber pressure over time, timestamps, machine states. These tools exist to decode that.

## Slide 3
Two tools, two jobs
UNanofabTools — DAT Tools
3
DATfixer — the translator
Reads a raw .DAT file.
Writes a clean, readable .txt file.
Can also draw a graph directly.
DATgrapher — the chart maker
Reads a cleaned .txt file.
Draws a pressure-vs-time graph.
Shows it on screen or saves an image.
Notes: Frame the pair simply. DATfixer is the translator: raw binary in, readable text out. DATgrapher is the chart maker: cleaned text in, a pressure-over-time graph out. You usually run DATfixer first, then DATgrapher on its output — though DATfixer can also produce a graph itself in one step. Translator then chart maker.

## Slide 4
The typical workflow
UNanofabTools — DAT Tools
4
1
Locate the .DAT file on the same Windows machine as the tools.
2
Run DATfixer translate the .DAT into a readable _cleaned.txt.
3
Read or graph open the text, or run DATgrapher on it for a chart.
Notes: Walk the normal path. Find the .DAT on the machine, run DATfixer to get a cleaned text file (named the same with '_cleaned' added), then either read that text directly or feed it to DATgrapher for a chart. Mention these are command-line tools run from PowerShell on Windows, on the same machine that holds the file — they're not part of the website.

## Slide 5
What's inside a .DAT file
UNanofabTools — DAT Tools
5
DATfixer recognizes the Denton's byte markers and decodes them:
raw bytes: 05 00 [8 bytes] 08 00 .. 00 ...
│ │ │
│ │ └─ field separator → a comma
│ └─ an 8-byte number → a pressure value
└─ 'a measurement follows' marker
becomes: 15:14:33), [0.000000486000], Auto, ...
Markers and raw numbers in → readable, comma-separated text out.
Notes: Give a peek under the hood without requiring code skills. DATfixer walks the file looking for the Denton's markers. A specific 2-byte marker means 'a measurement follows,' and the next 8 bytes are a number — the chamber pressure — which it writes in brackets. Another marker means 'separate these,' which it turns into a comma. It also strips the invisible junk characters and lines up timestamps. The result is a tidy, comma-separated, human-readable line. The exact byte values were reverse-engineered from the machine's output.

## Slide 6
DATfixer, step by step
UNanofabTools — DAT Tools
6
1
Decode measurements turn the byte markers into bracketed pressure numbers.
2
Insert separators turn delimiter markers into commas.
3
Strip the junk remove invisible control characters.
4
Tidy up put each timestamp on its own line; align the data.
5
Save write the result as a .txt next to the original.
Notes: This is the detailed 'what DATfixer does' slide. In order: decode the measurement markers into bracketed numbers; convert delimiter markers to commas; remove the invisible control characters that made it unreadable (keeping line breaks and tabs); tidy the layout so each timestamp starts a line and items are comma-separated; and finally save a .txt. The original .DAT is never modified.

## Slide 7
Running DATfixer
UNanofabTools — DAT Tools
7
Minimum: python.exe .\DATfixer.py YourLog.DAT
Option | What it does
-o NAME | Save the cleaned file somewhere else / under a new name
--graph | Also draw a pressure-vs-time graph as a .png
--strict | Remove even line breaks and tabs (denser, rarely needed)
--raw | Skip the 'make it pretty' formatting
--no-commas | Don't insert commas between items
-h / --help | Show the help message
Notes: Show the command and the options. The bare command produces YourLog_cleaned.txt. The most common addition is --graph, which also saves a PNG chart. The others are situational: -o to redirect output, --strict/--raw/--no-commas to change formatting. -h prints all of them. Don't read every row aloud; point out the bare command and --graph as the two people will use most.

## Slide 8
DATgrapher: charting the pressure
UNanofabTools — DAT Tools
8
Reads a cleaned text file and finds each timestamp and pressure.
Converts timestamps to 'seconds since the run started.'
Plots pressure over time — on screen by default.
A log scale option makes the huge range of vacuum pressures readable.
Notes: Explain DATgrapher. It scans a cleaned file for timestamps and bracketed pressure values, turns the times into seconds-from-start, and plots pressure over time, popping up a window by default. The --log option is genuinely useful: vacuum pressures span many orders of magnitude during a pump-down, and a logarithmic scale shows the whole curve instead of squashing the small numbers flat.

## Slide 9
Running DATgrapher
UNanofabTools — DAT Tools
9
Minimum: python.exe .\DATgrapher.py YourLog_cleaned.txt
Option | What it does
-o NAME | Save the graph to an image file
--no-display | Don't pop up a window; just save the image
--log | Use a logarithmic scale for the pressure axis
-h / --help | Show the help message
Notes: The DATgrapher command and options. Point it at the cleaned .txt, not the raw .DAT. By default it shows the graph; --no-display saves without a window (useful on a machine with no screen), and --log switches to the logarithmic pressure axis. Simpler than DATfixer — four options.

## Slide 10
Which tool when?
UNanofabTools — DAT Tools
10
Use DATfixer
To read a Denton log as text.
For a quick graph in one step (--graph).
Use DATgrapher
To plot an already-cleaned file.
To re-plot with a log scale.
Notes: A quick decision aid. Reach for DATfixer to turn a raw log into readable text, or to get a fast graph in a single command. Reach for DATgrapher when you already have cleaned files and want to chart them — especially to re-plot with a log scale. If in doubt: DATfixer first, DATgrapher second.

## Slide 11
Good to know (and a few rough edges)
UNanofabTools — DAT Tools
11
The tools never change your original .DAT — they only read it.
Point DATgrapher at the cleaned .txt, not the raw .DAT.
Runs longer than a day aren't fully handled (a midnight-only correction exists).
If only a few data points are found, you'll get a warning — usually a wrong-format input.
Notes: Set honest expectations. Originals are safe — read-only. A common mistake is pointing DATgrapher at the raw .DAT; use the cleaned text. The time handling assumes a run within a day with at most one midnight crossing. And a 'few data points' warning almost always means the input wasn't the expected format. These rough edges are documented in the developer notes for whoever maintains the tools.

## Slide 12
Recap
UNanofabTools — DAT Tools
12
Denton 635 logs are binary .DAT files — unreadable as-is.
DATfixer translates them into clean text (and can graph).
DATgrapher charts pressure vs. time from cleaned text.
Both are command-line tools run on Windows; originals stay untouched.
Notes: Wrap up. Two small, focused tools: one translates the Denton's binary logs to readable text, the other charts the pressure. Run them from PowerShell, originals are never modified. That's the whole story — a clean example of a small utility doing one job well. Open for questions.